## 查询结果
id | equipment_id | maintenance_order_id | order_number | maintenance_type | start_date | end_date | operator_id | operator | fault_description | maintenance_details | total_cost | remarks | created_at | updated_at | equipment_code | name | model | manufacturer | purchase_date | warranty_period | department | department_id | location | status | last_maintenance_date | next_maintenance_date | responsible_person | contact_number | purchase_price | service_life | description | image_url
12 | 12 | 12 | MO2025043000001 | repair | 2025-04-30 06:00 | 2025-05-01 08:45 | 18 | 杨主管 | 设备运行一段时间后自动关机，提示温度过高 | 冷却风扇故障，导致系统过热保护，更换冷却风扇组件，清理散热系统 | 4200.00 | 建议每月清理散热系统，确保通风良好 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | preventive | 2024-05-15 | 2024-05-16 09:30 | 18 | 杨主管 | 年度预防性维护 | 全面检查CT扫描仪各系统，清洁内部组件，更换滤网，校准X射线管，更新系统软件至V4.2.1 | 5800.00 | 年度维护完成，设备状态良好 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | preventive | 2024-08-12 01:00 | 2024-08-12 08:45 | 19 | 朱技术员 | 季度预防性维护 | 检查冷却系统，清洁散热器，测试各项功能参数，校准图像质量 | 2200.00 | 季度维护完成，冷却系统工作正常 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | preventive | 2024-11-20 00:30 | 2024-11-20 10:00 | 18 | 杨主管 | 季度预防性维护 | 检查电气系统，测试高压组件，清洁探测器，校准图像重建算法 | 2500.00 | 季度维护完成，建议更换备用电源电池 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | preventive | 2025-02-18 | 2025-02-19 04:30 | 19 | 朱技术员 | 季度预防性维护 | 检查机械部件，润滑转动组件，测试床位移动系统，校准激光定位器 | 2300.00 | 季度维护完成，更换了床位控制器的磨损零件 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | calibration | 2024-07-05 05:00 | 2024-07-05 09:30 | 18 | 杨主管 | 图像质量校准 | 使用标准模体进行图像质量校准，调整HU值，优化图像重建参数 | 1500.00 | 校准后图像质量提升15%，噪点减少 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | calibration | 2024-10-08 01:30 | 2024-10-08 07:00 | 19 | 朱技术员 | 剂量校准 | 使用剂量计进行X射线输出校准，调整mA和kV设置，优化剂量控制算法 | 1800.00 | 校准后患者辐射剂量降低10%，同时保持图像质量 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | calibration | 2025-01-12 02:00 | 2025-01-12 08:15 | 18 | 杨主管 | 系统校准 | 全面校准系统参数，包括探测器灵敏度、准直器精度、床位定位精度等 | 2000.00 | 校准完成，系统精度达到原厂标准 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | repair | 2024-06-23 02:15 | 2024-06-24 10:30 | 18 | 杨主管 | 控制台显示器故障，图像显示异常 | 更换显示器主板和电源模块，重新校准显示器色彩和亮度 | 3200.00 | 更换后显示效果恢复正常，建议使用稳压电源 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | repair | 2024-09-05 00:45 | 2024-09-06 04:30 | 19 | 朱技术员 | 床位移动卡顿，无法平滑运行 | 清理并润滑导轨，更换磨损的传动带和电机轴承 | 2800.00 | 维修后床位移动平稳，无噪音 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | repair | 2024-12-15 01:20 | 2024-12-16 06:45 | 18 | 杨主管 | 系统启动时报错，无法进入操作界面 | 修复系统引导文件，重装部分损坏的软件组件，恢复系统配置 | 1500.00 | 系统恢复正常，备份了系统配置以防再次出现问题 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg
12 | 12 |  |  | repair | 2025-03-10 | 2025-03-11 09:30 | 18 | 杨主管 | 系统软件升级 | 升级操作系统至最新版本，安装新的图像处理算法，更新安全补丁 | 3500.00 | 升级后系统性能提升20%，新增AI辅助诊断功能 | 2025-05-14 14:29 | 2025-05-14 14:29 | CT-001 | CT扫描仪 | uCT 530 | 上海联影医疗科技有限公司 | 2023-12-14 16:00 | 2028-12-14 16:00 | 放射科 | 16 | CT室F-102 | normal |  | 2025-12-14 16:00 | 徐志强 | 15612345678 | 3500000.00 | 15 | 64排螺旋CT，具有低剂量扫描和智能图像重建功能 | https://example.com/images/uct-530.jpg